# The kiosk is stuck in one of the pre-scan stages Impedance out of range issues

Issue: I am trying to do a scan and in one of the stages before I scan the process gets stuck and I cannot progress

### Troubleshooting: Kiosk Scan Not Registering

If the kiosk is not registering a scan, please follow these steps:

#### ✅ Pre-Scan Checklist

Ensure the following conditions are met before starting a scan:

- You are standing still on the kiosk.

- The kiosk is placed on a flat, solid surface.

- You are not wearing socks or shoes.

- Your hands are firmly touching the metal strips on the handles.

#### 🔍 If the Issue Persists

Try the following troubleshooting steps:

- Check Kiosk Connectivity
Ensure the kiosk is online and connected to the network.

- Capture the Issue
Take a photo of the screen or error where the kiosk is getting stuck and send it to the support team. Support team will look to see where the issue is whether the scale is working and got to the next to start scan and then failed, or whether the scale is not working. If the scale is not working, you will see 0.00kg when the machine is stood on.

Please also get a photo of the handles and wires and its common for wires to be exposed which in turn means the scale isn't working and the current is not going through

- Change Scale Mode
Switch the scale mode from Precision to Normal, or vice versa.

- Calibrate the scale through control panel settings- click on scale, then calibrate in top right-hand corner, then hit save

- Update the Software
Check for and install any available software updates.

- Try a different PowerPoint on a different side of the gym (different powerline)

- Escalate to Level 2 Support
If none of the above steps resolve the issue, escalate to Level 2 for further investigation or hardware servicing. If the issue persists after completing the above steps, escalate the case to Level 2 Support for further diagnosis and resolution. Please reply to customer saying that this has been escalated to level 2 and they will require new parts. In the email, please include whether customer is in warranty or not. In a separate issue forward the HubSpot ticket to level 2

  - Level 2 will communicate with regional manager of the required part(s) for the repair and with approval ship parts either directly to the gym or to the technician to complete the repair.

  - Once machine is repaired level 2 organise the return of the faulty parts. They then will see whether customer was in warranty or not. If they are in warranty,  Evolt will cover all charges, if they are out of warranty Evolt will pass on costs- parts, shipping and tech fees to customers.

  - Last step is to return used parts to Selvas for assessment. Parts can be sent every few months in one big order.