# Network API Status is Disconnecting

### Issue

EVOLT Scanner is stuck on the Control Panel and the Network API Status is Disconnecting

1. Connect to the Affected Machine via TeamViewer

- Launch TeamViewer on your local computer.

- Enter the TeamViewer ID of the target machine.

- When prompted, enter the password: EvoltG2o18.

- Wait until the remote session is fully established.

2. Open a Terminal Window

- On the remote machine, press Ctrl + Alt + T to open a terminal window.

3. Stop All Backend Services

- In the terminal, type the following command and press Enter:

  - sudo supervisorctl stop all

- Wait a few seconds for all services to stop before proceeding to the next step.

4. Check the Current System Time Configuration

- Type the following command and press Enter:

  - timedatectl status

- Review the output and confirm whether NTP synchronized is set to no.

5. Enable Network Time Protocol (NTP)

- To synchronize the system time, type the following command and press Enter:

  - sudo timedatectl set-ntp true

6. Restart All Backend Services

- Restart the previously stopped services by typing and press Enter:

  - sudo supervisorctl restart all

- Wait until you see confirmation that the services are now RUNNING.

7. Refresh the Control Panel Interface

- Go to the control panel interface in your browser or application.

- Click Refresh or reload the page.

- The Network status should now show as Connected.

- The “Start Scanning” button should be visible and operational.

Verification

- Monitor the control panel to ensure the Network API remains in the Connected state.

- Click the “Start Scanning” button to confirm it initiates scanning successfully.

- Review any logs or notifications to confirm communication between the front end and back end is stable.

Troubleshooting

If the issue persists:

- Verify internet connectivity on the target machine.

- Ensure there are no typos in the commands.

- Confirm the system time zone is correct.

- If necessary, escalate to a system administrator for deeper analysis.