# The Scale API Status is disconnecting; Scanner will not move past Control Panel

### Issue

If the Scale API Status is disconnected; then the scanner will not move past the Control Panel and there is no Red Start Scanning Button. You will also see the Scale Status is Unavailable.

- Ensure that the software is current and up to date on the scanner; if is not:

  - Run a Software Update ensuring that the customer does not power the unit off during the process.

- If after the update is completed and the software version is current; power cycle the unit.

- If after the power cycle the issue persists; ask the customer to test the unit on a different outlet as in some cases the power output can be related to the issue.

- If all of the above does not resolve the issue; a replacement of the AP PCB and/or Power PCB will be necessary; escalate to the EVOLT Team for arrangement.